## Slide 1
Development of a Dual Wavelength Microchip Laser for NASA’s Raman Mass Spectrometer (RAMS)
Matthew Mullin, Jane Lee, Tony Yu, Molly Fahey, Andrej Grubisic
AEROSPACE CONFERENCE 2025

## Slide 2: Intro
Introduction:
NASA Goddard Space Flight Center (GSFC)
Greenbelt, Maryland USA
Laser & Electro Optics Branch
Planetary Environments Lab
AEROSPACE CONFERENCE 2025

## Slide 3
Moon
Mars
Europa
Titan
We are building laser systems for planetary exploration platforms.
Science Objectives:
Search for the evidence of extraterrestrial life – “bio signatures”
Assess the habitability of environments beyond Earth.
Better understand the timing and formation of the Solar System.
Identify potentially viable economic resources for utilization.
Targets:
Enceladus
Images: NASA/JPL-Caltech
AEROSPACE CONFERENCE 2025

## Slide 4
4
AEROSPACE CONFERENCE 2025
The Raman Mass Spectrometer Instrument (RAMS)
+
Scanning
Chemical Maps of Samples
*These are simulations for visualization purposes
Raman
Mass Spec
=
Mass Spectrometry
Detector
Laser beam
Ion Flow
Raman Spectroscopy
+
Laser beam
Sample
Spectrometer
Filter

## Slide 5
5
AEROSPACE CONFERENCE 2025
Chemical Maps of Samples
Better understand extraterrestrial environments to assess the potential for life beyond Earth
The Raman Mass Spectrometer Instrument (RAMS)
Characterize Organic Carbon Reservoirs
Identify Minerals Present
Map Organic/Inorganic Composition

## Slide 6
Time of Flight Mass Spectrometer
RAMS Laser System
Raman Spectrometer
TVAC Chamber
Mass Spectrometer Collection Tube
Achromatic Focusing Lens
Translation Stage
Sample
Raman Collection Fiber Optics
MEMS Steering Mirror
RAMS Instrument Overview:
AEROSPACE CONFERENCE 2025

## Slide 7: The RAMS Laser System Requirements:
RAMS Laser Requirements | 
Wavelengths | 257.5 nm 515 nm
Fluence Range (variable) | UV: 0.02-1.0 J/cm2 VIS: 0.01-1.0 J/cm2
Spot Size (1/e2) | ≤ 10 µm
Shot-to-shot stability | ≤ 10 % (Goal: ≤ 5%)
Pulse Rate | 1-10 kHz (selectable) + Burst Capable
Spectral Bandwidth | ≤ 0.02 nm
Beam Steering | ≤ ± 6°
Pulse Width | ~ 1 ns
AEROSPACE CONFERENCE 2025
Need two wavelengths for two different science modes.
Ultraviolet
257.5 nm
Green
515 nm
Laser Desorption Mass Spectrometry
Raman Spectroscopy

## Slide 8: The RAMS Laser System Block Diagram:
10 w 940 nm Laser Diode
110 um Fiber Optic
Aspherical Pump Lenses and Back Reflection Filter
Yb:YAG & Cr:YAG Microchip Crystal
Volume Bragg Grating Output Coupler
Collimating Lens
1030 nm Half Wave Plate on Motorized Rotation Stage
515 nm Half Wave Plate
Beam Dump
515 nm Half Wave Plate #2
KTP Crystal
(515 nm Generation)
BBO Crystal
(257.5 nm Generation)
RTP High Voltage Electro-Optic
Polarizing Beam Splitting Cube
Wavelength Selection Filter
Output to MEMS Mirror
AEROSPACE CONFERENCE 2025

## Slide 9
The RAMS Laser System – As Built:

## Slide 10
Pump Diode (940 nm)
Microchip Laser Cavity (1030 nm)
1030 nm Waveplate
515 nm Waveplate #1
515 nm Waveplate #2
515 nm Filter on Motor Flip Mount
High Voltage Driver
High Voltage Pulse Picker Optic (RTP)
BBO UV Crystal
KTP Green Crystal
Start Pulse Detector
Collimating & Focusing Optics
Exit Aperture
Collimating Lens
The RAMS Laser System – As Built:
Green Gen Module
Pulse Picking Module
UV Generation Module

## Slide 11
Raman Mass Spectrometer Laser Demonstration
MEMS Mirror Scanning
Laser Operation at 515 nm at full power
AEROSPACE CONFERENCE 2025

## Slide 12: Fundamental Laser Cavity Performance (1030 nm):
Chart: Repetition Rate & Pulse Energy vs Current (1030 nm)
| 4 | 4.5 | 5 | 5.5 | 6 | 6.5 | 7 | 7.5 | 8 | 8.5 | 9 | 9.5 | 10 | 10.5 | 11 | 11.5
Repitition Rate (Hz) | 1534 | 2829 | 3602 | 3570 | 4682 | 5854 | 6848 | 7820 | 8409 | 7740 | 8800 | 9266 | 9650 | 10043 | 10100
Pulse Energy (uJ) | 116 | 131 | 145 | 171 | 163 | 171 | 179 | 186 | 192 | 208 | 195 | 199 | 202 | 205 | 208
1030 nm Collimated Far Field Beam
1030 nm, 1.4 ns Pulse
10 kHz Operation Demonstrated with 0.5% stability
Data taken from IR Cavity detector on breadboard at 8.5A output
AEROSPACE CONFERENCE 2025

## Slide 13: Optical Design: Non-Linear Conversion
13
AEROSPACE CONFERENCE 2025
1030 nm
515 nm
257.5 nm
Yb:YAG Laser Cavity
KTP Crystal
BBO Crystal
Potassium titanyl phosphate
Beta barium borate

## Slide 14: Green Conversion (515 nm) Stage Performance:
Chart: Repetition Rate & Pulse Energy vs Current (515 nm)
| 4 | 4.5 | 5 | 5.5 | 6 | 6.5 | 7 | 7.5 | 8 | 8.5 | 9 | 9.5 | 10 | 10.5 | 11
Repetition Rate (Hz) | 1534 | 2829 | 3602 | 3570 | 4682 | 5854 | 6848 | 7820 | 8409 | 7740 | 8800 | 9266 | 9650 | 10043 | 10100
Pulse Energy (uJ) | 33 | 41.5 | 43.6 | 46.5 | 78.4 | 80.6 | 84.2 | 84 | 91 | 91 | 100.8 | 100.3 | 99.1 | 99.8 | 101
515 nm Collimated Far Field Beam
515 nm, 1.3 ns Pulse
KTP Crystal
Motorized 1030 nm Stage
Laser Cavity Chiller Lines
Energy Meter
Microchip Laser Cavity
AEROSPACE CONFERENCE 2025

## Slide 15: Energy Throttling By Half Waveplate Rotation (515nm)
Chart: Green Channel Output Energy vs Waveplate Rotation Angle
| 80.5 | 80 | 79.5 | 79 | 78.8 | 78.599999999999994 | 78.400000000000006 | 78.099999999999994 | 77.900000000000006 | 77.599999999999994 | 77.37 | 77.099999999999994 | 76.8 | 76.569999999999993 | 76.3 | 76 | 75.7 | 75.48 | 75.2 | 74.8 | 74.599999999999994 | 73.8 | 73 | 72.7 | 72 | 71 | 70 | 69 | 68 | 67 | 66 | 65 | 64 | 63 | 62 | 61 | 60 | 59 | 58 | 57 | 56 | 55.7 | 55 | 54 | 53
Green Channel Waveplat atttenuatiom Energy output | 87 | 93.5 | 95.8 | 97 | 98.4 | 99.5 | 97.8 | 95.6 | 94.3 | 93.8 | 89.6 | 85.1 | 83.6 | 80.1 | 76.8 | 74.7 | 71.7 | 69.3 | 65 | 63 | 62 | 59 | 55.2 | 51 | 46.5 | 39.7 | 34.2 | 30.2 | 26.1 | 22.4 | 18.5 | 15.3 | 12.5 | 9.2 | 6.6 | 4.9 | 3.6 | 2.4 | 1.5 | 0.91 | 0.62 | 0.624 | 0.71 | 1.1 | 1.8
Raman Levels
Energy Throttling By Half Waveplate Rotation (515nm)
AEROSPACE CONFERENCE 2025

## Slide 16: RAMS Laser: Diode Current vs UV Output & Burst Control
Chart: Repetition Rate & Pulse Energy vs Current (257.5 nm)
| 4 | 4.5 | 5 | 5.5 | 6 | 6.5 | 7 | 7.5 | 8 | 8.5 | 9 | 9.5 | 10 | 10.5 | 11
Repitition Rate (Hz) | 2043 | 3140 | 3851 | 4090 | 4820 | 5990 | 7260 | 8100 | 8200 | 8250 | 8620 | 9420 | 9750 | 10100 | 10450
Pulse Energy (uJ) | 8.151 | 10.2505 | 10.7692 | 11.4855 | 19.3648 | 19.9082 | 20.7974 | 20.748 | 22.477 | 22.477 | 24.8976 | 24.7741 | 24.4777 | 24.6506 | 25.2
The 50 shot burst capability of the laser system achieved by firing the HV RTP optic. Each burst is fired every 2 seconds. The burst-to-burst average pulse energy remained within +/- 8% for the duration of testing.
Chart: RAMS Laser 257.5 nm Burst Capability
Focused UV Beam at ~50 um diameter
AEROSPACE CONFERENCE 2025

## Slide 17
RAMS 515 nm Laser Beam Shaping at Different Raster Positions
Normal Incidence Angle
4 mm across
*Beam images are zoomed in (gaps not to scale)
Demonstrating Rastering the RAMS Laser Across Sample Plane
MEMS Mirror Steering Demonstration
Green Beam on Sample
UV Beam on Sample
Green Beam on Sample
~ 15 um Diameter
~ 15 um Diameter

## Slide 18: Challenges we faced:
Laser induced damage in microchip cavity.
Thermal control of laser cavity.
Tiny Optics (cavity is about the size of a grain of rice)
Nonlinearity of wavelength conversion 2x & 4x’s any instability
Many Bells & Whistles = Complex Opto-mechanical systems
Tiny chiller plate prototype
Intense pump focusing lead to many damage events
Complex & Large setup for a “microchip laser”
AEROSPACE CONFERENCE 2025

## Slide 19
RAMS Laser Delivered To Mass Spectrometer Lab
Validation

## Slide 20
Ruggedization
Proto-flight version passed GEVS vibration and hermeticity Testing in FY24

## Slide 21
Optimization
FY25 R&D effort to miniaturize
And optimize for scale.
FY2025 IRAD:
Ultra-Compact Space Flight Lasers for Handheld, Rover, and Drone Applications

## Slide 22: Summary
We successfully demonstrated a dual channel (UV & Green) laser for the RAMS instrument meeting all requirements for testing.
The laser has been delivered to the RAMS instrument for testing – in progress.
We have demonstrated ruggedized version of laser assembly (GEVS Vibration).
We are pursuing rapid miniaturization of the technology.
AEROSPACE CONFERENCE 2025

## Slide 23
Thank You!
NASA Goddard Planetary Environments Lab
NASA Planetary Instrument Concepts for the Advancement of Solar System Observations (PICASSO) Office
IEEE Aerospace Conference